Step-by-step start-up procedure: **LSM 780\_FCC**

* Remove ***Dust Cover***
* With the ***Switches Panel***: follow these steps sequentially and wait few seconds between each step
  + Turn ON ***Main*** switch
  + Turn ON ***System PC***
  + Turn ON ***Components***
* If you require the Argon laser, follow these steps
  + Turn ON Main switch (can be left ON at all time)
  + Turn the key on the ON position (horizontal)
  + Wait 5 minutes
  + On the power controller, set the switch from ***idle*** to ***run*** position
  + When the green LED is ON, turn the knob until the red LED lights up, then decrease the power so that the red LED is off
* Switch ON the computer (this can be done while starting the Argon laser)
* If you require the fluorescent light source, switch ON the X-Cite lamp

*Note: the lamp should not be switched OFF (ON) within 25 minutes it has been switched ON (OFF)*