Step-by-step shut down procedure: **LSM 780\_FCC**

* Switch OFF the X-Cite lamp, if it was ON

*Note: the lamp should not be switched OFF (ON) within 25 minutes it has been switched ON (OFF)*

* Switch OFF all lasers via the software
* Quit ZEN
* Wait in Windows for 2 minutes (or until the hard disk is not turning, i.e no files being written)
* Quit Windows, and wait for Windows to shut down properly
* Switch OFF the Argon laser manually
	+ Switch the power controller from **run** to ***idle***
	+ Turn the laser key on the OFF position (vertical)
	+ Wait for fan to stop
* Switch off the computer
* With the ***Switches Panel***: follow these steps sequentially and wait few seconds between each step
	+ Turn OFF ***System PC***
	+ Turn OFF ***Components***
	+ Turn OFF ***Main*** switch
* Put ***Dust Cover*** back on the microscope